## Slide 1
1680
1930
1960
1515
Optical Microscopy

## Slide 2
E. coli
2 µm
Lymphocyte
7 µm
Viruses
0.02 -0.3 µm
Micro ---- "small"
Skopeîn --- "see"
Optical Microscopy
Pollen
40 µm
Human Hair
80 µm
Too micro to see
We can see

## Slide 3
F
Magnifying

## Slide 4
Magnifying

## Slide 5
Magnifying

## Slide 6
Lens

## Slide 7
  Chromatic aberrations arise from variations in the refractive indices of the wide range of frequencies in visible light (color fringes)
Chromatic aberrations

## Slide 8
Convex
Concave
Chromatic aberrations

## Slide 9
Aberrations are related to the spherical nature of the lens (edge blur)
Geometrical or Spherical aberrations

## Slide 10
Objectives
Leeuwenhoek 1680
Zeiss 2000
Lens
Triplet
Group
Lens
Triplet
Group
Lens
Doublet
Groups
Miniscus
Lens
Hemispherical
Front Lens
Planapochromat
$uper corrected

## Slide 11
Optical aberrations

## Slide 12
Immersion Oil
c - the speed of light in vacuum
v - the speed of light in the substance
Refractive Index n =
C
v
Glass n = 1.52
Air n = 1.000
Immersion Oil
n = 1.515

## Slide 13
Lighting

## Slide 14
Lighting
critical to quantitative microscopy
Neutral Density filter
Daylight balancing filter
CONDENSER

## Slide 15
Aplanatic-Achromatic
CONDEN$ER
Lighting
Condencer
Top
Lens
Internal
Lens
Elements
Aperture
Diaphragm

## Slide 16
Microscope
Eyepiece
Barrel
Focusing Screw
Objective
Hooke 1670
Olympus BX 2000
Cooling
Vents
Episcopic
Lamphouse
Lamp
Collector
Lens
Microscope
Control
Circuit
Board
Diascopic
Lamphouse
Line
Cord
Transformer
Collector
Lens
Focus
Knob
Filters
Mirror
Illumination
Intensity
Adjustement
Base
Field
Diaphragm
Condenser
Mecanical
Stage
Objective
Breathshield
Vertical
Illuminator
Dichroic
Filters
Beamsplitters
Eyepiece
(Ocular)
Rubber
Eyecup
Lenses
Nose-
piece
Extension
Tube
Shutter
Focusing
Telescope
Camera
Blck Box
Film Plan
Electrical Contacts
Automatic
Exposure
Monitor
Light
Sensor
Projection
Eyepiece
Trinocular
Head
Prisms
Shutter

## Slide 17
Color Fidelity

## Slide 18
Color Fidelity
CCD Detector

## Slide 19
Color Fidelity
Monitor Calibration
Calibration Slide

## Slide 20
$
$
$$
$
Optical aberration$$$$$
$$$$
$$$

## Slide 21
Optical aberration$$$$$
Computational compensation

## Slide 22
Computer compensated microscopy
jointly compensates
-- optical aberrations
-- illumination non-uniformities
-- imaging noise within a probabilistic framework.

## Slide 23
Computer compensated microscopy
Corrected Intensity Distributions using Regularized Energy minimization
CIDRE
K. Smith, Y. Li, F. Piccinini, G. Csucs, C. Balazs, A. Bevilacqua, and P. Horvath, “Cidre: an illumination-correction method for optical microscopy,” Nature methods, vol. 12, no. 5, pp. 404– 406, 2015.
As is
CIDREd

## Slide 24
F
Magnifying

## Slide 25
F
Magnifying

## Slide 26
F
Magnifying

## Slide 30
Absorption
Fluorescence
Polarization
Refraction
Phase shift
Optical Microscopy

## Slide 31
Brightfield
Phase contrast
DIC
Fluorescent
Absorption
Fluorescence
Polarization
Refraction
Phase shift
Optical Microscopy

## Slide 32
400 nm
Fluorescence
JABLONSKI DIAGRAM
450 nm

## Slide 33
Incoming
photon
K
L
M
N
Fluorescence
Fluorescence

## Slide 34
Fluorescence
Mucicarmine stain

## Slide 35
Stain -Tissue interaction
Tissue
+
_
Stain
Ionic
bonds
+
_
Hydrogen
bonds
Van der Waals
forces
Covalent
bonds
electron delocalization
FRET
Fluorescence-lifetime
changes
Fluorescence spectra modulations
Quenching
characterize the molecular microenvironment adjacent to ...

## Slide 36
Fluorescence
Mucicarmine stain
Local Discriminant Bases
Nearest-Neighbor classification
Spectral classifier
Hierarchical fusion using vector quantization
Unmixing-based constrained nonnegative matrix factorization
Spectral Weighting Envelopes
1BT Based Band Selection
Spectral Angle Mapper
Complete Linear Spectral Unmixing
Minimum Noise Fraction Transformation
Principal Component Transformation

## Slide 37
Hyperspectral
Notes: Journal of Biomedical Optics 19(1), 010901 (January 2014)
Medical hyperspectral imaging: a review

## Slide 38
Specimen Thickness

## Slide 39
Point Spread Function
Z
Working distance

## Slide 40
Point Spread Function
PSF describes the imaging system response to a point input
Z
In microscopy the point spread functions is asymmetric due to lens imperfections

## Slide 41
Deconvolution
Optical sectioning
Fast iterative soft-thresholding
Richardson–Lucy total variation
thresholded Landweber
DeconvolutionLab2: An open-source software for deconvolution microscopy.
https://www.ncbi.nlm.nih.gov/pubmed/28057586
Fast iterative soft-thresholding

## Slide 42
Manz at al. * 2000
AutoDeblur from Media Cybernetics
Deconvolution
*J Microbiol Methods 2000, 40:125-34.
Tissue sections of Chondrosia reniformis after hybridization with the Bacteria specific probe EUB338

## Slide 43
Focus

## Slide 45
Dr. Marvin Minsky
Confocal microscope: 1957
Confocal microscope

## Slide 46
LASER
Light
Amplification by the
Stimulated
Emission of
Radiation
Scanning
Confocal
Microscope
Laser
Confocal microscope

## Slide 47
Dichroic
Mirror
Confocal microscope

## Slide 48
Conventional / Confocal

## Slide 49
Conventional / Confocal

## Slide 50
Light Sheet Microscopy
Single Plane Illumination Microscopy (SPIM)
Medaka juveniles
Japanese rice fish
Philipp Keller, Stelzer Group, EMBL

## Slide 51
Biplan Localization Microscopy
+ 400 nm
- 400 nm
0
Cylindrical
lens

## Slide 52
Biplan Localization Microscopy
Vutara, Inc..

## Slide 53
F
Magnifying

## Slide 54
Conventional / Confocal

## Slide 55
Confocal

## Slide 56
)))))))))))))
)))))))))))))
))))))))
)))))))))))))
Diffraction

## Slide 57
Airy disk
Diffraction

## Slide 58
Resolution
?
?

## Slide 59
Resolution
Abbe diffraction limit
Abbe Resolutionx,y = λ/2NA
Numerical Aperture
NA = n•sin(θ)
n - refractive index of the imaging medium ( air, oil)
θ - aperture angle (1,4 in the best case)
D = 200 nm

## Slide 60
Resolution
Magic wand

## Slide 61
Expansion microscopy
Expansion microscopy: principles and uses in biological research
Asmamaw T. Wassie, Yongxin Zhao & Edward S. Boyden
Nature Methods volume 16, pages 33–41 (2019)
Swellable Polyelectrolyte
Hydrogel

## Slide 62
Expansion microscopy
Expansion microscopy: principles and uses in biological research
Asmamaw T. Wassie, Yongxin Zhao & Edward S. Boyden
Nature Methods volume 16, pages 33–41 (2019)
Escherichia coli immunolabeled for membrane lipopolysaccharides
up to ≈ 20 nm ?

## Slide 63
Deconvolution
Values from one function are removed from the values of another.
Naive inverse filtering
Tikhonov regularization
Regularized inverse filtering
Landweber
Tikhonov–Miller
Fast iterative soft-thresholding
Richardson–Lucy
Richardson–Lucy with total-variation regularization

## Slide 64
Deconvolution
Leica-microsystems
up to 100 nm

## Slide 65
Near-field optical microscopy

## Slide 66
Near-field optical microscopy
special ($)1.78 refractive index coverslips
special ($ $) 1.78 refractive index oil
special ($ $ $) objective 100x 1.65NA

## Slide 67
Near-field optical microscopy

## Slide 68
Total internal reflection
Θ1
Θ2
The critical angle is the angle of incidence above which the total internal reflectance occurs
Evanescent wave
≈100 nm
up to 20 nm
Evanescent waves quantum tunneling phenomenon
Product of Schrödinger wave equations

## Slide 69
Total Internal Reflection Fluorescence Microscopy (TIRF)
McKinney et al.Nature Methods 6, 131 - 133 (2009)
5 µm
1 µm
up to 20 nm

## Slide 70
Structured Illumination Microscopy
The Moire effect
+
=
-
=
?
!

## Slide 71
Structured Illumination Microscopy

## Slide 72
Structured Illumination Microscopy
up to 100 nm

## Slide 73
Switch to nonfluorescente state
400 nm
Triplet state
Non fluorescent
Fluorescence

## Slide 74
PA-GFP
Emission
Desactivation
Emission
Desactivation
Super-resolution
Optical Fluctuation Imaging (SOFI)

## Slide 75
Super-resolution
Optical Fluctuation Imaging (SOFI)
Fast, background-free, 3D super-resolution optical fluctuation imaging (SOFI).
Dertinger T, Colyer R, Iyer G, Weiss S, Enderlein J.
Proc Natl Acad Sci U S A. 2009
t1
:
:
:
:
tn
t1
:
:
:
:
tn

## Slide 76
Super-resolution
Optical Fluctuation Imaging (SOFI)
Fast, background-free, 3D super-resolution optical fluctuation imaging (SOFI).
Dertinger T, Colyer R, Iyer G, Weiss S, Enderlein J.
Proc Natl Acad Sci U S A. 2009
Richardson–Lucy total variation
Richardson–Lucy total variation
up to 50 nm

## Slide 77
doughnut-shape
+
=
Stimulated emission depletion
488 nm
520 nm
Stefan Walter Hell

## Slide 78
Stimulated emission depletion
(STED)
up to 50 nm

## Slide 79
Fluorescence Localization Microscopy

## Slide 80
PA-GFP
PA-GFP
Fluorescence Photoactivation Localization
Microscopy
PA-GFP
Emission
Localization
(calculation)
Photobleaching
Activation

## Slide 81
Fluorescence Localization Microscopy

## Slide 82
Fluorescence Photoactivation Localization
Microscopy (PALM)
Hess, S.T., T.P. Girirajan, and M.D. Mason. 2006. Ultra-high resolution imaging by fluorescence photoactivation localization microscopy. Biophys J. 91(11):

## Slide 83
ZHUANG LAB/HARVARD UNIV.
Fluorescence Photoactivation Localization
Microscopy
up to 20 nm

## Slide 84
Super-Resolution Microscopy
Fluorescence nanoscopy in cell biology.
Sahl SJ, Hell SW, Jakobs S.
Nat Rev Mol Cell Biol. 2017 Nov;18(11):685-701.

## Slide 85
Super-Resolution Microscopy
Super-resolution microscopy demystified.
Schermelleh L, Ferrand A, Huser T, Eggeling C, Sauer M, Biehlmaier O, Drummen GPC.
Nat Cell Biol. 2019 Jan;21(1):72-84. doi: 10.1038/s41556-018-0251-8. Epub 2019 Jan 2. Review.

## Slide 86
Merci !
Merci !

## Slide 87
http://zeiss-campus.magnet.fsu.edu/index.html
http://www.olympusmicro.com/primer/anatomy/anatomy.html
https://www.microscopyu.com/microscopy-basics

## Slide 88
400 nm
Fluorescence
Jablonski
diagram
450 nm
JABLONSKI DIAGRAM

## Slide 89
400 nm
Fluorescence (true)
Jablonski
diagram
450 nm
JABLONSKI DIAGRAM

## Slide 90
FL1
F
FL2
FL1
FRET
1-10 nm
Time (ns)
Fluorescence-Lifetime Imaging (FLIM)
Free
Coupled

## Slide 91
Fluorescence-Lifetime Imaging (FLIM)
Lin HJ et al. Cytometry A. 2003
Fluorescence lifetime-resolved pH imaging of living cells.
Alpy F et al. J Cell Sci. 2013
STARD3 or STARD3NL and VAP form a novel molecular tether between late endosomes and the ER.
Molecular Interactions
Intracellular pH

## Slide 92
Z Compensate System
Confocal Olympus FV1000
543 nm
633 nm
453 nm
488 nm
515 nm
Z
Y
X
l
Time
Fluorescence
- Lifetime
Hexadimensional
(space- time)
Microscopy
Hooke 1670

## Slide 93
Confocal PSF
WF
CF

## Slide 94
400 nm
Fluorescence
Jablonski
diagram
450 nm
JABLONSKI DIAGRAM

## Slide 95
Wavelength and Light Properties
575 nm
E = h c /l
(l)
High
Energy
Low
Energy
0,5* *2

## Slide 96
Multiphoton exitation
800 nm
800 nm
1200 nm
1200 nm
1200 nm
400 nm
Dr. Maria Göppert-Mayer : theory of two-photon quantum transitions
(two-photon absorption and emission) 1931,
Jablonski
diagram
450 nm
JABLONSKI DIAGRAM

## Slide 97
LASER
500 nm
Pulsed LASER
1000 nm
1200 nm
800 nm
nm
500
nm
800
1200
Temporal photon concentration
Wavelength Tunability

## Slide 98
Prof. Watt W. Webb et al.
Two-photon laser scanning
fluorescence microscopy: 1990
Spatial photon concentration

## Slide 99
Confocal / Multiphoton

## Slide 100
Luo at al. Cell Structure and Function 2006, 31: 63
Comparison of photoactivation of PA-GFP in vivo with single-photon (405 nm) and multiphoton (790 nm) laser light.
Photoconversion
Excitation area

## Slide 101
Thorsten at al. Nature 2004; 427: 159
T-cell priming by dendritic cells in lymph nodes occurs in three distinct phases
In VIVO Imaging

## Slide 102
Conventional / Confocal / Biphoton

## Slide 103
Science Vol. 385, No. 6713 ; 6 Sep 2024

## Slide 104
Anisotropic optical properties of molecules
Multiphoton polarization microscopy
polarizator
“Biphotonic”
Laser
Linear dichroism

## Slide 105
Biphoton polarization microscopy
Cell expressing GAP43-CFP-Gαi2 and α2a-adrenergic receptor
G-proteins orientation
+ Norepinephrine
Base line
Lazar J et al. Nat Methods. 2013
Two-photon polarization microscopy reveals protein structure and function

## Slide 106
YFP
YFP
CFP
CFP
FRET
1-10 nm
Förster Resonance Energy Transfer (FRET)
Molecular Basis

## Slide 107
Z Compensate System
Confocal Olympus FV1000
543 nm
633 nm
453 nm
488 nm
515 nm
Z
Y
X
l
Time
Fluorescence
- Lifetime
Hexadimensional
(space- time)
Microscopy
Hooke 1670

## Slide 108
Super-resolution
Bleaching Assisted Localization Microscopy(BALM)
t1
:
:
:
:
tn
t1
:
:
:
:
tn
Fast, Bleaching/blinking assisted localization microscopy for
superresolution imaging using standard fluorescent molecules.
Burnette DT, Sengupta P, Dai Y, Lippincott-Schwartz J, Kachar B.
Proc Natl Acad Sci U S A. 2011 Dec 27;108(52):21081-6

## Slide 109
Up to 50 nm of lateral resolution and ? nm of axial resolution
Super-resolution
Bleaching Assisted Localization Microscopy(BALM)
Fast, Bleaching/blinking assisted localization microscopy for
superresolution imaging using standard fluorescent molecules.
Burnette DT, Sengupta P, Dai Y, Lippincott-Schwartz J, Kachar B.
Proc Natl Acad Sci U S A. 2011 Dec 27;108(52):21081-6

## Slide 110
+
=
X Y
X Y Z
Stimulated emission depletion 3X

## Slide 111
Stimulated emission depletion
(STED 3X)
Up to 50 nm of lateral and axial resolution

## Slide 112
Switch to nonfluorescente state
550 nm
Triplet state
Non fluorescent
Fluorescence
600 nm
400 nm
Thiols
(R-SH)
Ground state depletion
Ground state
JABLONSKI DIAGRAM

## Slide 113
FITC
FITC
Ground State Depletion Microscopy
direct Stochastic Optical Reconstruction Microscopy
FITC
Stochastic Activation
Localization
(calculation)
Emission
Total Deactivation
(Ground State Depletion)

## Slide 114
Ground State Depletion Microscopy

## Slide 115
Optical Microscopy
Brightfield
Phase contrast
DIC
Fluorescent
Absorption
Fluorescence
Polarization
Refraction
Phase shift

## Slide 116
Light-Tissue interaction
RAMAN
AutoFluorescence
Absorbance
Fluorescence
Stain
Fixation (or not)
RAMAN: https://en.wikipedia.org/wiki/Raman_scattering
Wave shape
Changes
(𝜆 ; polarization)